| SECTION B: Renal and GIT |  |  |
| --- | --- | --- |
| Item | ZNB No. | Page No. |
| - Bag urine drainage |  | 1 |
| - Bag urine collection - paed adhesive | 6085 | 1 |
| - Bag Hanger for urine drainage bag |  | 1 |
| - Bottles for pasteurizer |  | 1 |
| - Bottle Lids for pasteurizer |  | 2 |
| - Catheters-Urinary | 6085 | 2 |
| - Latex |  | 2 |
| - Silicon-no bulb |  | 3 |
| - Silicon with bulb |  | 3 |
| - Cup-feeding Sinappi | RT9/2020 | 4 |
| - Feeding tubes | 6685 | 4 |
| - Riles |  | 4 |
| - Jejunal with introducer |  | 4 |
| - Jejunal-without introducer |  | 5 |
| - Nasogastric |  | 5 |
| - Kangaroo feeding pump bags |  | 6 |
| - Napkins | 5760 | 7 |
| - Peritoneal Dialysis | 6730 | 8 |
| - pH test Strips | 6745 | 8 |
| - Litmus paper-blue | 6745 | 8 |
| - Stoma bags |  | 9 |

1. 30 01861 Bag urine collection - paed adhesive

- Bag, Urine Sample Collecting, Paediatric. Polyethylene with non-allergenic adhesive face plate. Unisex. - Size: 55 mm x 150 mm - Sterile, individually peel packed - PACK: Box of 10

2. 30 018 60 Bag-Urine drainage

- In terms of the latest issue of CKS 463. Plasticised, medical grade PVC, the finish of which prevents adhesion between front and back of bag. Type 2 - double ended with bottom drain Grade S; - With Urine sampling sleeve - With outlet tap and with non-return valve at upper end. - Sterile, individually peel packed, Method of sterilisation and expiry date to be reflected on the exterior of the pack.

Bag-Urine drainage: Paediatric: rated volume of ?ml

Bag-Urine drainage Neonatal: rated volume of ?ml

3. 30 317 04 Bag hanger -Urine drainage

4. Bottles-Pasteuriser

A. 30 047 00 Bottle- Pasteurisation and Human Milk Storage-Glass

Purpose: Reusable glass storage bottles suitable for the pasteurisation of human donor milk - Volume: 125 ml - Neck closure: 48mm single-thread - Material: Glass - The bottle must be supplied with a suitable lid - Heat resistant and autoclavable up to 120 °C - Packaging: Units

B. 30 04702 Bottle-Pasteurisation and Human Milk Storage-Plastic

Purpose: Reusable plastic storage bottles suitable for the storage, pasteurisation and freezing of human donor milk. - Volume: 130 ml - Neck Diameter: 50 mm - Material: Clarified bisphenol A-free polypropylene - The bottles must be - Made from transparent, shatter-proof material that is clearly graduated - Heat resistant and autoclavable up to 120 °C - Sterile and delivered in a tamper evident bags. - Packaging: Units

5. Bottle Lids-Pasteuriser

A. 30 047 00 Lid-glass

Purpose: Lid to seal pasteurisation jars and prevent air ingress or leakage - Dimensions: 48mm - Lid: single-thread - Material: Metal with Plastisol

B. 30 04702 Lids with foil seal for Human Milk Storage Bottles - Reusable

Purpose: Reusable screw cap lids with foil seal to provide secure sealing of BPA-free Pasteurisation Bottles - Consists of a plastic screw cap that fits securely on the BPA-free plastic bottle - The screw cap must have a foil seal to provide secure sealing during pasteurisation, and transportation up to the point of use - No water should leak into the human milk during human milk during pasteurisation and cooling - The seal should have a tab to avoid touching the neck of the bottle when removing the tamper evident foil seal. - Heat resistant and autoclavable up to 120 °C - Packaging: Units

6. Catheter urinary

A. Latex

COLLECTIVE REQUIRMENTS - Chlorinated surface to protect against protein allergy. - Distal end: One tapered end and one balloon port with a non-return valve and fitting for a syringe. - Proximal end: Closed round tip with TWO lateral eyes distal. - CAPACITY OF BALLOON. Minimum of 3 ml (proximal end) - Sterile, individually peel-packed, Method of sterilisation and expiry date to be reflected on the exterior of the pack. - With inner sleeve. - TO COMPLY WITH THE LATEST ISSUE OF SABS EN 1616 - PACK: BOX OF 10

31 032 98 Catheter, urethral, foley, two-way latex. 6 FG (Outer Diameter mm). See COLLECTIVE REQUIREMENTS Catheter urinary- Latex

31 032 99 Catheter, urethral, foley, two-way latex. 8 FG (Outer Diameter mm). See COLLECTIVE REQUIREMENTS Catheter urinary- Latex

31 032 00 Catheter, urethral, foley, two-way latex. 10 FG (Outer Diameter mm). See COLLECTIVE REQUIREMENTS Catheter urinary- Latex

31 032 01 Catheter, urethral, foley, two-way latex. 12 FG (Outer Diameter mm). See COLLECTIVE REQUIREMENTS Catheter urinary- Latex

B. Catheter, urethral, foley, two-way, silicone without introducer or bulb.

- Distal end: one tapered end - Proximal end: Closed round tip with TWO lateral eyes - Manufactured from 100% silicone - Sterile - Individual peel pack with inner sleeve. - The following must be noted on the packaging: - Trade name of product - Size and specification - Method of sterilization - Manufacturing site - CE number - Lot number - Expiry date PACK: BOX OF 10

C. Catheter, urethral, foley, two-way, silicone with introducer and bulb.

COLLECTIVE REQUIREMENTS - Distal end: one tapered end and one balloon port with a non-return valve and fitting for syringe - Proximal end: Closed round tip with TWO lateral eyes with retention balloon. - Manufactured from 100% silicone - Sterile - Individual peel pack with inner sleeve. - The following must be noted on the packaging: - Trade name of product - Size and specification - Method of sterilization - Manufacturing site - CE number - Lot number - Expiry date - PACK: BOX OF 10

31 040 18 Catheter, urethral, foley, two-way, silicone with introducer and bulb. 6FG (Outer Diameter 2.0 mm). Bulb capacity: 1.5ml-3 ml See COLLECTIVE REQUIREMENTS Catheter, urethral, foley, two-way, silicone with introducer and bulb.

31 040 19 Catheter, urethral, foley, two-way, silicone with introducer and bulb. 8FG (Outer Diameter 2.6 mm). Bulb capacity: 3ml See COLLECTIVE REQUIREMENTS Catheter, urethral, foley, two-way, silicone with introducer and bulb.

31 040 20 Catheter, urethral, foley, two-way, silicone with introducer and bulb. 10 FG (Outer Diameter 3 mm). Bulb capacity: 5ml See COLLECTIVE REQUIREMENTS Catheter, urethral, foley, two-way, silicone with introducer and bulb.

31 040 21 Catheter, urethral, foley, two-way, silicone with introducer and bulb. 12 FG (Outer Diameter 4 mm). Bulb capacity: 5ml See COLLECTIVE REQUIREMENTS Catheter, urethral, foley, two-way, silicone with introducer and bulb.

8. 30 208 00 Cup- feeding Sinappi

- Virgin high-grade polypropylene with stand temperature of 132 degree Celsius. - 200ml main holder - Calibrations marked every 25 ml - Additional feeding cup 50ml - 60ml - No sharp edges - Bbpa free

9. Feeding tubes

A. Nasogastric tube duodenal surgical ryles

COLLECTIVE REQUIREMENTS Purpose: Feeding in babies who are unable to take orally - Radio-opaque line. - Markings at 450 mm, 550 mm, 650 mm and 750 mm from proximal tip. - Distal end: Funnel connector with Luer syringe adaptor and captive plug. - Proximal end: Weighted, closed round tip with FOUR lateral eyes. - Manufactured from: medical grade, non-toxic PVC. - Sterile ,Individual peel pack - Method of sterilisation and expiry date to be reflected on the exterior of the pack. - PACK: BOX OF 10 See COLLECTIVE REQUIREMENTS Nasogastric tube duodenal surgical ryles

33 006 79 Nasogastric tube duodenal surgical ryles 8FG (Outer Diameter 4 mm). Minimum length 90 cm See COLLECTIVE REQUIREMENTS Nasogastric tube duodenal surgical ryles

33 006 80 Nasogastric tube duodenal surgical ryles 10FG (Outer Diameter mm). Minimum length 90 cm.

33 006 81 Nasogastric tube duodenal surgical ryles 12FG (Outer Diameter mm). Minimum length 130 cm. See COLLECTIVE REQUIREMENTS Nasogastric tube duodenal surgical ryles

33 006 82 Nasogastric tube duodenal surgical ryles 14FG (Outer Diameter mm). Minimum length 130 cm. See COLLECTIVE REQUIREMENTS Nasogastric tube duodenal surgical ryles

B. Jejunal tube with introducer

COLLECTIVE REQUIREMENTS Purpose: Naso-jejunal feeding of infants and children - Catheter must be soft, kink resistant, radiopaque silicone. - Must be graduated in centimeters - Open distal end with 2 lateral eyes and atraumatic tip - Must have introducer and weighted end for easy placement - Proximal end must fit a luer syringe and have a plug to close the end. - Manufactured from silicone and medical grade plastic - Latex free - Sterile and individually packed in peel pouch - Single use only - The following must be noted on the packaging: - Trade name of product - Size and specification - Method of sterilization - Manufacturing site - CE number - Lot number - Expiry date - Per unit

33 006 69 Jejunal tube Size: 6 Fr Length : 550 mm +/- 10% See COLLECTIVE REQUIREMENTS Jejunal tube with introducer

33 006 70 Jejunal tube Size: 8 Fr Length : 550 mm +/- 10% See COLLECTIVE REQUIREMENTS Jejunal tube with introducer

33 006 71 Jejunal tube Size: 10 Fr Length : 550 mm +/- 10% See COLLECTIVE REQUIREMENTS Jejunal tube with introducer

C. Jejunal tube without introducer and weighted tip 5FG

Purpose: Naso-jejunal feeding of neonates - Catheter must be soft, kink resistant, radiopaque silicone. - Must be graduated in centimetres - Open distal end with 2 lateral eyes and atraumatic tip - Proximal end must fit a luer syringe and have a plug to close the end. - Manufactured from silicone and medical grade plastic - Latex free - Sterile and individually packed in peel pouch - Single use only - The following must be noted on the packaging: - Trade name of product - Size and specification - Method of sterilization - Manufacturing site - CE number - Lot number - Expiry date Per unit

D. Nasogastric tube

COLLECTIVE REQUIREMENTS Purpose: Naso- or Orogastric feeding of infants and children - Catheter must be soft, kink resistant pvc and have a radiopaque line. - Must be graduated 20cms from distal end - Closed distal end with 2 lateral eyes and atraumatic tip - Proximal end colour code and must fit a luer syringe and have a plug to close the end. - Manufactured from pvc and medical grade plastic - Latex free - Sterile and individually packed in peel pouch - Must be packed straight, so catheter does not develop memory - Single use only - The following must be noted on the packaging: - Trade name of product - Size and specification - Method of sterilization - Manufacturing site - CE number - Lot number - Expiry date - Per unit

33 006 61 Nasogastric tube Size: 5 Fr Length: 400 mm +/- 10% Colour: Gray See COLLECTIVE REQUIREMENTS Nasogastric tube

33 006 63 Nasogastric tube Size: 6 Fr Length: 400 mm +/- 10% Colour: Light green See COLLECTIVE REQUIREMENTS Nasogastric tube

33 006 65 Nasogastric tube Size: 8 Fr Length: 400 mm +/- 10% Colour: Blue See COLLECTIVE REQUIREMENTS Nasogastric tube

33 006 68 Nasogastric tube Size: 10 Fr Length: 600 mm +/- 10% Colour: Black See COLLECTIVE REQUIREMENTS Nasogastric tube

10. Kangaroo feeding pump bags

Bag- sterile 500ml - cap closure - ice-pouch - to fit enteral feeding pumps - calibrated units 100ml - bpa free - single use

Bag-sterile1000ml - cap closure - ice-pouch - to fit enteral feeding pumps - calibrated units 100ml - bpa free - single use

Bag-feeding flushing 1000ml - two bags for flushing and feeding - calibrated - cap closure - fits feeding pump - bpa free - sterile - single use

Bag-feeding flushing-universal - oral and catheter tip syringes, - feeding spike - flushing and feeding through one line and pump - 1000ml - bpa free - sterile - single use - per set

11. Napkins

COLLECTIVE REQUIREMENTS Purpose: Collection and retention of urine and stool - Curved, flexible, uniform anatomical shape - Fluid impervious outer layer, non woven inner lining and absorbant pad - Ultra absorbant –locks away urine and runny stools - Absorbent pad: - Must be centralized - Must be free from contamination(eg adhesives/ manufacturing materials/hard spots) - Absorbent filler must be evenly distributed within the pad and must not be accessible from outside the napkin - Breathable to facilitate dry skin - Correctly placed elastic gathers around legs with anti-leak guard - Refastenable anywhere tapes - Must be sealed around the edges - Napkins must be folded flat without excessive crumpling - Must not be joined together due to an outflow of adhesive - Must comply with SANS 9001 quality standards - Packaging must bear the following legible and indelible information - Manufacturers name and / or trade mark - Size - Batch or date coding

Nappies Extreme Preterm <1000g - Designed specifically for extremely preterm babies - Width between legs +- 3cm See COLLECTIVE REQUIREMENTS Napkins

Nappies Preterm 1000-2000g - Designed specifically for preterm babies - Width between legs +- 5cm - PACKAGING: +-30 diapers per pack See COLLECTIVE REQUIREMENTS Napkins

08 001 08 Nappies Term 2500-4000g - Designed specifically for term babies - Width between legs +- 8cm - Super absorbent mass (SAM) 4g/ms, total diaper mass +/-28,4g/ms. PACKAGING: +-68 diapers per pack (4 packets per bale) See COLLECTIVE REQUIREMENTS Napkins

08 001 10 Nappies Infant- medium 5-6kg - End fold 106mm, max fold with width 120mm, - Super absorbent mass (SAM) 5g/ms, total diaper mass 34,4g/ms. PACKAGING: 200 each per case See COLLECTIVE REQUIREMENTS Napkins

08 001 12 Nappies Infant- large 7-12kg - End fold 106mm, max fold with width 120mm - Super absorbent mass (SAM) 6g/ms, total diaper mass 40.4g/ms. PACKAGING: 200 each per case See COLLECTIVE REQUIREMENTS Napkins

12. Peritoneal Dialysis (PD)

- 30 018 41 Drainage Bag

- 30 066 02 Peritoneal Dialysis Cap

- 30 072 15 Peritoneal Dialysis Clamp

- 30 062 13/ 30 490 19 Peritoneal Dialysis Giving Set

13. 30 445 25 pH test Strips

Purpose: To indicate if nasogastric tubes are in the stomach - Colour coded pH indicator strips with visual comparison guide on container. - CE certified for in vitro testing of pH in human gastric aspirate. - Clear graduations between pH steps particularly 4.5-5.0 and 5.0-5.5 - Packaging: Unit (1 tub with 150 strips)

14. 30 445 23 Litmus Paper-Dark blue

Purpose: To test whether a solution is acidic or basic. - Paper must turn red under acidic conditions - Colour change occurring over the pH range 4.5–8.3 at 25 °C - Tearable paper strips +- 0.5cm x 5cm - 20 strips bound in a booklet - Packaging: Unit ( 1 pack with 5 booklets-100 strips)

15. 3379949-74 Stoma bags

33 79949 Stomahesive Wafer Flexible purple 100mm - 125163 (15.2cm x 15.2cm)

33 79953 Stomahesive Wafer Orange 38mm - 125143

33 79958 Stomahesive Flexible Wafer Red 57mm - 125145

33 79960 Stoma - Urostome Bag Red 57mm - 400991

33 79963 Stoma - Pouch Urostomy Transparent Bag Orange 38mm – 400989

33 79967 Stoma - Pouch Drainable Standard Transparent Orange 38mm - 401601

33 79971 Stoma - Pouch Drainable Standard Opaque 57mm - 402534

33 79974 Stoma - Drainage Bag Purple 100mm - 401608