| SECTION E: MISCELLANEOUS |  |  |
| --- | --- | --- |
| Item | ZNB No. | Page No. |
| - Adaptor plug-catheter large | 6745 | 1 |
| - Applicator-Orange stick/haemolysis | 6745 | 1 |
| - Band Personal Identification | 6745 | 2 |
| - Caps -baby fabric |  | 2 |
| - Collar and cuff | 6745 | 2 |
| - Combs-nit/loose (fine-toothed double-edged comb) |  | 2 |
| - Cover-reflective temperature probe (critter cover) |  | 2 |
| - Developmental care products |  | 3 |
| - Bendable positioning aid |  | 3 |
| - Nest |  | 3 |
| - Gel positioning aid |  | 3 |
| - Dummies (Pacifiers)-Nonnutritive sucking devices |  | 3 |
| - ECG electrodes | 6745 | 4 |
| - Forceps |  | 4 |
| - Halstead Mosquito |  | 4 |
| - Iris |  | 4 |
| - Galley pot |  | 5 |
| - Gel-ultrasound | 6745 | 5 |
| - Kangaroo mother care items. |  | 6 |
| - Baby beanie-Preterm |  | 6 |
| - Baby beanie-Term |  | 6 |
| - KMC wrap |  | 6 |
| - KMC lounger |  | 6 |
| - Medicine measure |  | 6 |
| - Pad sanitary maternity |  | 6 |
| - Phototherapy eye shields |  | 7 |
| - Reusable |  | 7 |
| - Disposable |  | 7 |
| - Pipe Cleaners | 6745 | 7 |
| - Probe -Button end |  | 8 |
| - Razor disposable double edge | 6745 | 8 |
| - Receiver-small plastic |  | 8 |
| - Safety Pins |  | 8 |
| - Sheath holder-medial (sats probe grip) |  | 8 |
| - Spatula wooden tongue depressor | 6745 | 8 |
| - Stethoscope | 6745 | 8 |
| - Storage bin-plastic | 6745 | 9 |
| - Straws -Drinking | 6745 | 9 |
| - Tape measure | 6745 | 9 |
| - Thermometers digital underarm | 6745 | 9 |
| - Traction | 6745 | 9 |
| - Umbilical securement |  | 10 |

- 30 534 82 Adaptor plug-catheter large

- Spigot, Plastic, autoclavable. Tapered from specified diameter to point. - Diameter: 10 mm. - Length: ±52 mm. - PACK: 10 per box.

- Applicators

30 010 12 orange stick - Made from round smooth hardwood - Diameter:2 mm Length:150 mm - Autoclavable - Packaging: Box of 100

Applicator haemolysis

- 30 028 10-11 Bands- Personal Identification

COLLECTIVE REQUIREMENTS Purpose: Patient identification - Manufactured from plastic or other suitable material that is: - extra-soft, pliable and durable in order to prevent abrasion of delicate skin - hypoallergenic - Band width 2cm to allow for printing, in an indelible manner, the patient's name and other particulars - Band to be supplied with a locking device which cannot ordinarily be opened after first closure - Plastic press studs must fasten with ease. - Packaging: Box of 100

30 028 10 Child - Colour: White - Locking device from: 9cm - 21cm - Printing on box to read: Child/ Paediatric - See Band- Personal identification- COLLECTIVE REQUIREMENTS

30 028 11 Neonate - Colour: White - Locking device from: 6cm - 12cm - Printing on box to read: Neonates. - See Band- Personal identification- COLLECTIVE REQUIREMENTS

- Caps -baby fabric

- Size: 0-3months - Length: +-12cm - Width:+-7cm - Fabric: 2-Ply soft cotton-Non-toxic, skin-friendly, and breathable - Machine washable - Per unit

- 30 079 03 Collar and cuff

- Made from soft foam covered with cotton woven mesh - Height and width: 50 mm x 10 mm. - Length: 4 m. - To include a minimum of 5 cable ties. - PACK: Roll

- Combs-nit/loose (fine toothed double edged comb)

Purpose: Removal of lice, nits and eggs from hair - Sturdy plastic handle - Stainless steel teeth-fine toothed - Rounded teeth to prevent lacerating the skin - Easy to clean - Weight: +-36g - 100% integrity after boiling

- Cover-reflective temperature probe

Purpose: Attachment of temperature probe to neonatal skin and refection of and protection from ambient and radiant heat sources in order to obtain accurate skin temperature readings. - Gentle, non-irritating hydrocolloid or foam adhesive (not hydrogel) - Does not contain natural rubber latex or plasticizer DEHP - Highly reflective gold mylar or foil upper surface/cover - Per 100 units

- Developmental care products

- Bendable positioning aid Purpose: To facilitate developmentally supportive positioning in sick and preterm babies. - Moldable lead rod wrapped in thick spongy foam. - Removable washable cover - Per unit

Small: L 71cm x H 6.4cm

Large: L 91.4 cm x H 8.9 cm.

- Fabric positioning nest Purpose: To facilitate developmentally supportive positioning in sick and preterm babies. - Reusable and washable - Fabric: Cotton or Flannel - Padded foot roll - Adjustable straps x2 - Per unit

Small- to fit baby <1000g

Medium- to fit baby 1000g 2500g

Large- to fit baby >2500g

- Gel positioning aid Purpose: Pressure relieving device to prevent pressure areas - Round - Soft flexible outer shell - Filled with nontoxic gel and antibacterial agent - CE certified - Latex and BPHA free - Per unit

Small: 1cm thick

Large: 1.5cm thick

- Dummies (Pacifiers)-Nonnutritive sucking devices

COLLECTIVE REQUIREMENTS Purpose: Provision of non-nutritive sucking for sick or preterm babies - Designed for hospital use only - Designed to prevent nipple confusion. - Made from one solid piece of medical-grade soft silicone - Must be phthalate, BPA and latex-free - Easy to clean and can be sterilized - Has a big, soft perioral guard (mouth shield) to stimulate nerve endings around the mouth - Has a nose cut out that allows space for oral and nasal tubing - The nipple - Is cylindrically/symmetrically shaped to make sure the tongue’s pressure against the palate is evenly distributed (tongue cupping) - Base is narrow so baby can easily close their mouth around it. (similar to thumb or mom’s nipple) - Has a small bulge at the distal end to stimulate limbic system of brain - An optimal length to reach the ridge between soft and hard palate - Packaged in sealed plastic pouch/bag - Packaging: Unit

- Term Dummy Size: Term babies Should be similar to the size and shape of mother’s nipple See: Nonnutritive sucking devices- COLLECTIVE REQUIREMENTS

- Preterm Dummy Size: Preterm babies <2kg. Should be similar to the size and shape of preterm thumb with smaller perioral guard. See: Nonnutritive sucking devices -COLLECTIVE REQUIREMENTS

- 30 248 04-30 248 37 ECG Electrodes

COLLECTIVE REQUIREMENTS Purpose: For diagnostic and monitoring purpose of an ECG tracing. - The electrode must be easy to strip from the backing liner - The sponge backing materials must protect the sensor and gel from surgical and cleaning fluids and still permit maximum flexibility - The adhesive backing must provide good contact and resist electrode contact loss due to perspiration. - There must be a non-adhesive edge on one side of the backing to facilitate removal. - The electrode must be disposable using standard waste disposal procedures - Must have a shelf-life of 2 years if unopened - Must be SABS 1461 or CE approved - Must be hypoallergenic and latex free - Packaging to clearly state: - Type and size of electrode - SABS/CE compliance - Expiry date - Packaging: Pack of 30

30 248 04 Solid Gel ECG Electrode -Paediatric Consists of a pre-gelled Silver/Silver-Chloride solid gel, spongeless electrode with a press stud that fits standard ECG leads Electrode size: < 50mm wide See: Dummies (Pacifiers)-Nonnutritive sucking devices COLLECTIVE REQUIREMENTS

30 248 09 Sponge ECG Electrode – Paediatric Consists of a pre-gelled Silver/Silver-Chloride self-adhering sponge electrode with a press stud that fits standard ECG leads Electrode size: < 50mm wide See: Dummies (Pacifiers)-Nonnutritive sucking devices COLLECTIVE REQUIREMENTS

30 248 27 Tape ECG Electrode – Paediatric Consists of a pre-gelled Silver/Silver-Chloride self-adhering tape/cloth electrode with a press stud that fits standard ECG leads Electrode size: < 50mm wide See: Dummies (Pacifiers)-Nonnutritive sucking devices COLLECTIVE REQUIREMENTS

30 248 37 / 3025026 Tape ECG Electrode – Neonate Consists of a pre-gelled Silver/Silver-Chloride self-adhering tape/cloth electrode with a press stud that fits standard ECG leads Electrode size: ≤ 30mm wide See: Dummies (Pacifiers)-Nonnutritive sucking devices COLLECTIVE REQUIREMENTS

- Forceps

COLLECTIVE REQUIREMENTS Purpose: For maintaining sterility during procedures and when fine control is required. Must be: - Non sterile - Reusable - Durable high grade surgical stainless-steel - Satin Steel Finishing - Boil Tested, Performance Tested, Shape Tested - Autoclavable at 134 C - No chrome plating - Adheres to international ISO/CE standards - Bidder to stipulate rust protection procedure - Has a 10-year guarantee

- Forceps- Halstead Mosquito - Length: +/- 125 mm - Profile: Straight - Finger ring handle - Serrated jaw - Includes a locking mechanism (ratchet) - Ref. no. BH 11OR See: Forceps COLLECTIVE REQUIREMENTS

- Forceps- Iris

Straight-toothed - Profile: Straight - Length: +-10cm - Fine tip width: +-0.8 m - Jaw: 1:2 teeth. Teeth length: +-1.4mm - With guide-pin and guide hole. - Ribbed grip - Ref. no AE 586-10 See: Forceps COLLECTIVE REQUIREMENTS

Non toothed - Profile: Straight - Length: +-10cm - Fine tip width: +- 0.8 m - With guide-pin and guide hole. - Serrated jaw - Ribbed grip - Ref. no AE 270-10 See: Forceps COLLECTIVE REQUIREMENTS

- Galley pot

Purpose: To contain fluids during a sterile procedure - Size: +- 10cm x5cm - Volume: +- 200ml - Small round cup shaped bowl - Non sterile - Reusable - Durable high grade surgical stainless-steel - Satin Steel Finishing - Boil Tested, Performance Tested, Shape Tested - Autoclavable at 134 C - No chrome plating - Adheres to international ISO/CE standards - Bidder to stipulate rust protection procedure - Has a 10-year guarantee

- 30 344 30 / 3021305 /3034429 Gel-ultrasound

Purpose: To facilitate ultrasound imaging - Hypoallergenic and non-toxic - Water soluble, non-staining and easily cleanable - Does not contain any lipid components - Free from formaldehyde and salt - Odorless - Gamma radiated to assure sterility - Will not damage the probe - Does not contain air bubbles - pH= 7 - 200-250ml bottle with dispensing nozzle - Per bottle

- Kangaroo mother care (KMC) products

- Baby beanie hat-term Purpose: To prevent heat loss from the head of a newborn baby. - To fit preterm babies between 28-36 weeks gestation. - Approximate head circumference 25-30cm. Should be able to fold up to fit different head sizes. - Made from hospital grade, latex free, machine washable, breathable, cotton rib-knit fabric. - Colour should be a pastel baby shade. - Packaging: 50 units in plastic packaging /box.

- Baby beanie hat-term Purpose: To prevent heat loss from the head of a newborn baby. - To fit term babies between 37- 42 weeks gestation. - Approximate head circumference 33-38 cm. Should be able to fold up to fit different head sizes. - Made from hospital grade, latex free, machine washable, breathable, cotton rib-knit fabric. - Colour should be a pastel baby shade. - Packaging: 50 units in plastic packaging /box.

- KMC wrap Purpose: To help parents provide continuous and quality skin-to-skin care to their newborns. The ergonomic wrap design keeps the baby securely in place, preventing occlusion of the airway, on the caregivers chest with little strain on the shoulders and back of the parent. - Designed to be able to adjust to accommodate both preterm and term babies. - A variety of designs will be considered but must meet the following requirements: - Made of soft but strong non stretch cotton, reusable after machine washing - Easy-to-use: Does not require an assistant to put on - Gender-neutral: Acceptable to wear by fathers as well as mothers - Designed for easy breastfeeding without removing newborn from thermal care - Allows unrestricted movement and causes minimal fatigue in the back and shoulders of the carrier - There are no knots in the back in order to ensure it is comfortable for the carrier to use while sitting or lying down - Safe for the baby: Keeps the baby secure in the correct position for providing effective skin-to-skin care - Packaging: 50 units in plastic packaging/box. - The following design is included for reference:

- Medicine measure liquid plastic cup

Purpose: The dispensing of medicine or pills - Material: medical grade polymer plastic (PP) - Colour: transparent - Volume: 30 mL - Calibre: +- 3.8 cm - Height: +- 3.5 cm - Graduated in millilitre increments - Scale is clearly marked for good visibility. Does not wipe off - The edge/lip of the cup is smooth. - CE Marked - Non sterile

- Pads- sanitary maternity

3044003 Pads Sanitary Regular With Loops Autoclavable Pkt Of 12

3044004 Pads Sanitary Maternity Self Adhesive Pkt Of 12

- Phototherapy eye shields

A. Reusable

COLLECTIVE REQUIRMENTS Purpose: Neonatal eye protection during UV light treatment in cases of jaundice Consists of two parts: strap and eye pads/goggles. Strap/headband must: - Be designed to prevent slipping and produce a secure fit, but be easy to move or adjust when necessary. - Be made of soft, slip resistant material that expands at the back to cup and grip the head even when the baby moves. - Have two ends which are temporarily mutually attachable by means of an adjustable fabric fastener. - Stretching of the strap should not result in stretching of the eye shield/goggles. Eye shield /goggles must: - Include ocular pockets designed for patient comfort and safety. - Seal around the skeletal socket reducing pressure and allowing the eyes to open and close naturally. - Be made from washable soft, hypo-allergenic, non-woven, light-blocking (including UV rays up to 450 Nm), and liquid-absorptive material. - Be shaped to accommodate the nose without the need to cut the fabric.

Small: To fit a head circumference +-20-25 cm See: Phototherapy eye shields-COLLECTIVE REQUIREMENTS

Medium: To fit a head circumference +- 26-34 cm See: Phototherapy eye shields-COLLECTIVE REQUIREMENTS

Large: To fit a head circumference +- 33-38 cm See: Phototherapy eye shields-COLLECTIVE REQUIREMENTS

B. Single use (Disposable)

COLLECTIVE REQUIREMENTS Purpose: Neonatal eye protection during UV light treatment in cases of jaundice - Made of antiallergenic, latex free, synthetic, stretch fibres that absorb perspiration from the skin, allowing faster evaporation - Has UV light blocking properties (up to 450 Nm) - Consists of a wraparound design with one or two adjustable attachment points - Must firmly grip the head preventing dislodgement when the baby moves - Has large ergonomic eye pads that limit pressure on the eye and prevents UV light leakage.

Small: To fit a head circumference +-20-25 cm See: Phototherapy eye shields-COLLECTIVE REQUIREMENTS

Medium: To fit a head circumference +- 26-34 cm See: Phototherapy eye shields-COLLECTIVE REQUIREMENTS

Large: To fit a head circumference +- 35-40cm See: Phototherapy eye shields-COLLECTIVE REQUIREMENTS

- Pipe Cleaners

- Probe -Button end

Purpose: For dilation of umbilical blood vessels - Non sterile - Reusable - Durable high grade surgical stainless-steel - Satin Steel Finishing - Boil Tested, Performance Tested, Shape Tested - Autoclavable at 134 C - No chrome plating - Adheres to international ISO/CE standards - Bidder to stipulate rust protection procedure - Has a 10-year guarantee - Rounded tip - Length: +-11-15cm - Tip width: +-1-1.5mm - Ref. no AW 090-11/BN 104R - Per unit

- 30 465 10 Razor disposable double edge

Purpose: For surgical shaving - Sharp, double edged with carbon steel or medical steel blade. - For single patient use. - Packaging: Box 100 pcs

- Receiver-small plastic

- Safety Pins

- Sheath holder-medical (sats probe grip)

Purpose: For securement of pulse oximeter “Y” saturation probes in neonates without damaging delicate skin - Soft, foam, tapeless reusable wrap that does not stick to the skin - Allows for easy application or removal - Secured with Velcro - Holds sensor firmly to prevent misalignment of optical components, while blocking out excessive ambient light - May also be cut to fit any length - +- 3cm x 14cm x 0.5cm - Per unit

- 30 045 70 Spatula wooden tongue depressor

Purpose: For depressing the tongue during an oral examination - Made from hardwood - Plastic is not acceptable - Must have rounded edges on either side - Sides and surfaces must be smooth but not slippery - Width: ± 17 mm Length: ±150 mm Thickness: ±1.5 mm - Autoclavable - Single use - Packaging: Box of 100

- Stethoscope

30 542 07 Paediatric stethoscope Purpose: To identify, listen to, and study heart, lung, and other body sounds for physical assessment and diagnosis - Consisting of a dual-sided chest piece with ± 3.3 cm diaphragm - The chest piece must provide high acoustic sensitivity for both high and low frequency sounds with both a floating diaphragm and open bell. - Tubing must be flexible, resilient and phthalate plasticizers free, ± 70 cm long - Headset tension must be adjustable for various head size - The ear tips must provide a good acoustic seal and comfortable fit. - The ear tubes must be angled to align with the ear canals and must be able to rotate - All components must be latex free - 5-year warranty must be included - Packaging: Unit

3054205 Stethoscope earpieces

- 74 80 024 -25 Storage bin –Plastic

COLLECTIVE REQUIREMENTS Purpose: Storage of stationary, pharmaceutical or surgical sundries. - Manufactured from thick robust plastic - Must be stackable - Front façade must allow access to content and space for a label - Must be available in different colours – red; blue yellow - State brand name - Per unit

74 80 024 Storage bin –Plastic Size 4 - Approximate Length: 380 mm - Width: 205 mm - Height: 170 mm

74 80 025 Storage bin –Plastic Size 5 - Approximate Length: 380 mm - Width: 420 mm - Height: 170 mm

- 3055601 Straws - Medicinal drinking

- Tape measure

Purpose: Measurement of length and head circumference in neonates - In hospital use - Plastic coated durable woven material tape - Measures in centimeters and inches - Durable markings that don’t wash off in alcohol - Length: 1.5m - Per unit

- 30 244 00 Thermometer -digital underarm

Purpose: Single patient use thermometers to record oral and axillary temperatures - Must be made of BPA free plastic with a flexible tip. Must have a storage case - Must be capable of giving fast accurate reading within 10 seconds - Must cover temperature ranges of at least: 34 °C to 42°C. (Not Fahrenheit) - Must have a memory to record the last temperature reading - Must have automatic shut off function - For oral and axillary use. - Must be completely sealed and compatible with alcohol disinfectant - Must be supplied with a battery - Per unit

- Traction

A. Kit set

B. Nylon Cord

- Umbilical securement

30 072 01 Clamp Purpose: To securely clamp the umbilical cord at birth - Made from plastic (polymer) - Single use - ± 5cm long - Finger grip to ensure safe and easy handling. - Clamp must be constructed for easy application in wet, slippery conditions. - Serrations and stoppers must prevent clamp from slipping but must not cut through the cord. - Safe security lock, with a click to indicate correct locking, to protect against accidental re-opening after clamping. - Individual sterile peel pack. - Gamma – sterilized. CKS 312 - The following must be noted on the packaging: - Trade name - Size and specification - Manufacturing site - CE number - Lot number - Expiry date - Per box of 100

30 578 27 Umbilical tape, Cotton-Sterile Purpose: Ligature thread for tying the umbilical cord. - Non-absorbable. - Braided and flat. - Material: 100 % cotton. Non fraying - Sterile. - Width : 3,0 mm - Length: 25,0 cm - The following must be noted on the packaging: - Trade name - Size and specification - Manufacturing site - CE number - Lot number - Expiry date - Per box of 12 sachets

Umbilical tape, Cotton-Nonsterile Purpose: Ligature thread for tying the umbilical cord. - Non-absorbable. - Braided and flat. - Material: 100 % cotton. Non fraying - Width: 3mm. - Length: 50m. - Single use. - Non-sterile - Labelling on the primary packaging: - Name and/or trademark of the manufacturer - Manufacturer's product reference - Lot number - Per roll